{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["MPE_Calculator","Temp_Comp_Calculator","LookupTables"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"MPE_Calculator","index":0,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[[8,4,8,5],[9,4,9,5],[9,7,9,8],[9,1,9,3],[1,0,1,9],[14,0,14,4]]}
[]
["MPEE0 Linear Distance Calculator"]
[]
[]
[null,{"$f":"IF(H9=\"inch\",G9,G9 / 25.4)","v":12},"inch",null,"=",null,{"$f":"IF(H9=\"mm\",G9,G9 * 25.4)","v":304.79999999999995},"mm"]
[]
[]
[]
[null,null,null,null,"Distance (L) ",null,12,"inch"]
[null,"Machine Specification      From Manufacturer",null,null,null,null,null,"Choose Unit of Measure From the Dropdown Menu."]
["MPEE0 = +/-",1.4,"+  L  /",333]
[]
[]
[]
["Maximum Allowable Error at Distance (L) = ",null,null,null,null,null,{"$f":"IF(H9=\"mm\",FIXED(((B11+(G5/D11))/1000),7),FIXED((((B11+(G5/D11))/1000)/25.4),7))&IF(H9=\"mm\",\" mm\",\" inch\")","v":"0.0000912 inch"}]
[null,null,null,null,null,null,{"$f":"IF(H9=\"inch\",FIXED(((B11+(G5/D11))/1000),7),FIXED((((B11+(G5/D11))/1000)/25.4),7))&IF(H9=\"inch\",\" mm\",\" inch\")","v":"0.0023153 mm"}]

{"sheet":"Temp_Comp_Calculator","index":1,"n_rows":22,"n_cols":9,"rows_kept":22,"cut":false,"hidden":false,"merged":[[21,5,21,6],[21,0,21,4],[19,6,19,8],[4,0,4,9],[1,0,1,9],[14,1,14,3],[20,0,20,4],[14,4,14,6],[20,5,20,6]]}
[]
["Temperature Compensation Calculator"]
[]
[]
["Thermal Compensation Calculator"]
[]
[]
[null,{"$f":"IF(I15=\"celsius\",H15 /( 5/9) + 32,H15)","v":71.349997999999999},"fahrenheit","=",null,null,{"$f":"IF(I15=\"celsius\",H15,(H15 - 32) *(5/9) )","v":21.86111},"celsius"]
[]
[null,{"$f":"IF(I17=\"inch\",H17,H17 / 25.4)","v":6.9989999999999997},"inch","=",null,null,{"$f":"IF(I17=\"mm\",H17,H17 * 25.4)","v":177.77459999999999},"mm"]
[]
[null,{"$f":"IF(RIGHT(C16, 1)=\"F\",B16,B16*(5/9))","v":6.3},"µin / inch / ° ","=",null,null,{"$f":"IF(RIGHT(C16, 1)=\"F\",B16/(5/9),B16)","v":11.34},"µm / meter / ° C"]
[]
[]
[null,"Specify Coefficient of Thermal Expansion",null,null,"Actual Temperature",null,null,21.86111,"celsius"]
[null,11.34,"µm / meter / °C"]
[null,null,null,null,null,null,"Distance",177.77459999999999,"mm"]
[]
[]
[null,null,null,null,null,null,"NIST suggested Uncertainty from Approximation of CTE"]
[{"$f":"\"Required Compensation Amount at \"&H17&\" \"&I17&\"'s\"","v":"Required Compensation Amount at 177.7746 mm's"},null,null,null,null,{"$f":"FIXED((68-B8) * B10 * (B12 / 1000000),7)&\" inch\"","v":"-0.0001477 inch"},null,"+ / -",{"$f":"FIXED(((68-B5) * B10 * (B12 / 1000000))*0.1,7)&\" inch\"","v":"0.0002998 inch"}]
[{"$f":"\"at \"&H15&\" degrees \"&I15","v":"at 21.86111 degrees celsius"},null,null,null,null,{"$f":"FIXED((20-G8) * G10 * (G12 / 1000000),7)&\" mm\"","v":"-0.0037519 mm"},null,0.1,{"$f":"FIXED(((20-G5) * G10 * (G12 / 1000000))*0.1,7)&\" mm\"","v":"0.0040319 mm"}]

{"sheet":"LookupTables","index":2,"n_rows":2,"n_cols":3,"rows_kept":2,"cut":false,"hidden":false,"merged":[]}
["inch","fahrenheit","µin / inch / °F"]
["mm","celsius","µm / meter / °C"]