|  |  | Joystick calibration. The CMM should be turned on with a connection between Calypso and the CMM. Y-axis calibration: On the joystick box press the (shift lower return, and then press the up arrow key (this will light up the (y) axis, then press the return key one time, then hold the y axis joystick in the neg position and press the return key, then hold the y axis joystick in the pos position and press the return key, then with the y axis joystick in the neutral position press the return key. X-axis calibration: Next: on the joystick box press the (shift lower return, and then press the right arrow key (this will light up the (x) axis, then press the return key one time, then hold the x axis joystick in the neg position and press the return key, then hold the x axis joystick in the pos position and press the return key, then with the x axis joystick in the neutral position press the return key. Z-axis calibration: Next: on the joystick box press the (shift lower return, and then press the down arrow key (this will light up the (z) axis, then press the return key one time, then hold the z axis joystick in the neg position and press the return key, then hold the z axis joystick in the pos position and press the return key, then with the z axis joystick in the neutral position press the return key. Rotary Table -axis calibration: Next: on the joystick box press the (shift lower return, and then press the left arrow key (this will light up the axis, then press the return key one time, then hold the z axis joystick in the left position and press the return key, then hold the z axis joystick in the right position and press the return key, then with the z axis joystick in the neutral position press the return key. Note this is for a rotary table but go through the calibration anyway because it is in the controller that it looks at your calibration. This should have your joystick calibrated, but when you turn the drives back on if it moves slowly in any direction you missed one of the steps, so try the calibration again and make sure you follow the steps. Also if you are still having problems with the joystick box most likely you will need to replace the box. If the joystick box is out of warranty and the RED shock detection indicator is RED, then there should be no harm in taking it apart and checking for loose connection, broken parts etc… BUT……. IF it is under WARRANTY… then leave it to the professionals. |
| --- | --- | --- |